## Sheet1
Parameters for DST Horizontal Echelle Spectrograph
2006-06-28
SPINOR: TI TC245
Wavelength | 8542 | Å |  | l
Slit width | 50 | microns |  | b
Focal length | 3040 | mm |  | L
Lines/mm | 308.5714 | 0.0032407 | 5 | 1/d | d
Spectral Pixel size | 8.5 | microns |  | p
Spatial Pixel size | 19.75 | microns
Grating length | 206 | mm |  | W
Blaze angle | 52 | deg
Incident Angle | 56.34 | deg |  | i
Order | 6 |  |  | m
slit length | 10 | mm |  | 2*h
Image scale at slit | 133 | mm/arc-sec |  |  |  | 7.8233
Slit spatial sample | 0.37594 | arc-sec
Detector spatial pixels | 230 | pixels
Detector spatial length | 4.5425 | mm
Detector spectral pixels | 256 | pixels
Detector spectral length | 2.176 | mm
Chosen f/# | 36
Projected Grating Width | 126.8263 | mm |  | pw=Wcos(blaze)
Geometrical Width | 84.4444 |  |  | gw=fl/f#
Diffraction Width | 103.8707 |  |  | dw=2*l/b |  | 49.8 | 50.6
Illumination Width | 133.8656 |  |  | iw=sqrt(gw^2+dw^2)
Effective Width | 126.8263 |  |  | lesser of iw and pw |  | 48.3 | 46.77 | 51.7
Required Grating Length | 205.9549
Required Grating Height | 94.4444
Reflection angle | 48.5167 | deg | -7.8233 | q = asin ( m l /d - sin(i) ) |  | 7321.7143 | 10250.4
Slit Magnification | 0.83675 |  |  | M=cos(i)/cos(q) |  | 6406.5 | 12813
Resolving Power | 370703.4436 |  |  | R = W(sin(i)+sin(q)) /l |  | not correct need eff.
Camera Lens fl | 1219 | mm |  | l
Dispersion | 3407133.9324 | m/m |  | lmd/cos(q)
| 0.34071 | microns/mÅ |  | D
Diffraction Resolution | 23.0427 | mÅ |  | l/R
Spectral pixel width | 24.9477 | mÅ/pixel |  | p/D | l
Spectral slit width | 49.2386 | mÅ |  | l*b*M/DL
Spectral Resolution | 59.8146 | mÅ |  | RSS
Spectrum length | 6.3866 | Å
Spectrum curvature | 3.4745e-06 | radians |  | sec(q) m l a^2 / 2d
| 10.5624 | microns |  | where a = h/L
| 0.21125 | pixels |  |  |  |  | pixels
Spatial coverage | 70.1376 | arc-sec | 85.1752 | cube size | 8.0115 | mm | 186.566
Spatial Pixel size | 0.37033 | arc-sec
Grating Efficiency | 0.99474 |  |  | i = sin(x)/x
Shading | 0.87708 |  |  | where x = p m [cos(blaze) -
Total | 0.87247 |  |  | sin(blaze)/tan((i+q)/2)]
0.1778
shade = 1. - sin(q - blaze)/sin(90. - blaze)
|  |  |  | 0.87708 | incidence
|  |  |  | 0.90131 | reflection
0.87708