## Sheet1
Parameters for DST Horizontal Echelle Spectrograph
2003-11-01
SPINOR:1280 x 1024
Wavelength | 6302 | Å |  | l |  | 236.8 | arc sec
Slit width | 40 | microns |  | b |  | 31.4944
Focal length | 3040 | mm |  | L |  | 22.2733 | cube size
Lines/mm | 308.5714 | 0.0032407 | 5 | 1/d | d
Spectral Pixel size | 12 | microns |  | p
Spatial Pixel size | 12 | microns
Grating length | 206 | mm |  | W
Blaze angle | 63.46 | deg
Incident Angle | 63.46 | deg |  | i
Order | 9 |  |  | m
slit length | 10 | mm |  | 2*h
Image scale at slit | 133 | mm/arc-sec
Slit spatial sample | 0.30075 | arc-sec
Detector spatial pixels | 640 | pixels
Detector spatial length | 7.68 | mm
Detector spectral pixels | 256 | pixels
Detector spectral length | 3.072 | mm
Chosen f/# | 36
Projected Grating Width | 92.0454 | mm |  | pw=Wcos(blaze)
Geometrical Width | 84.4444 |  |  | gw=fl/f#
Diffraction Width | 95.7904 |  |  | dw=2*l/b
Illumination Width | 127.6976 |  |  | iw=sqrt(gw^2+dw^2)
Effective Width | 92.0454 |  |  | lesser of iw and pw
Required Grating Length | 214.2917
Required Grating Height | 94.4444
Reflection angle | 58.8186 | deg | -4.6414 | q = asin ( m l /d - sin(i) ) |  | 5671.8 | 7089.75
Slit Magnification | 0.86301 |  |  | M=cos(i)/cos(q) |  | 5156.1818 | 8102.5714
Resolving Power | 572091.3756 |  |  | R = W(sin(i)+sin(q)) /l |  | not correct need eff.
Camera Lens fl | 1000 | mm |  | l
Dispersion | 5363879.2931 | m/m |  | lmd/cos(q)
| 0.53639 | microns/mÅ |  | D
Diffraction Resolution | 11.0157 | mÅ |  | l/R
Spectral pixel width | 22.3719 | mÅ/pixel |  | p/D | l
Spectral slit width | 21.1701 | mÅ |  | l*b*M/DL
Spectral Resolution | 32.7112 | mÅ |  | RSS
Spectrum length | 5.7272 | Å
Spectrum curvature | 5.2979e-06 | radians |  | sec(q) m l a^2 / 2d
| 16.1056 | microns |  | where a = h/L
| 0.40264 | pixels |  |  |  |  | pixels
Spatial coverage | 160.5053 | arc-sec | 175.5429 | cube size | 16.5115 | mm | 533.68
Spatial Pixel size | 0.27429 | arc-sec
Grating Efficiency | 0.73855 |  |  | i = sin(x)/x
Shading | 0.8189 |  |  | where x = p m [cos(blaze) -
Total | 0.6048 |  |  | sin(blaze)/tan((i+q)/2)]
-1.3073
shade = 1. - sin(q - blaze)/sin(90. - blaze)
|  |  |  | 1 | incidence
|  |  |  | 0.8189 | reflection
0.8189